SIMULATION OF AN HPGe DETECTOR AND A MARINELLI’S BEAKER SOURCE TO DERIVE EFFICIENCY CURVES WITH FLUKA 2005.6

A.A. Porta, F. Campi

Nuclear Engineering department – CeSNEF, Polytechnic of Milan

### Short informations

In this document you can find:

- a description of the work done with methods and results involving a Marinelli’s beaker certified calibration source as a multiline reference

- the layouts of .inp and source.f files used as FLUKA input and custom-source respectively

- the HPGe layout as computed by SimpleGeo beta version 1.0 (Chris Theis and Karl Heinz Buchegger)

### Methods

FLUKA code was used in its lower-limit energy range with a custom source to simulate photons emission and transport from a Marinelli’s beaker multiline calibration source to an HPGe detector. This way it is possible to construct a topic total-efficiency curve without having a specific calibration reference.

The first step foresees to know as well as possible the inner geometry of the HPGe endcup. The task is quite simple if it possible to make a radiography or a gammagraphy of the device.

In figure 1 it is possible to see the result of a gammagraphy made with Cs-137 of such a detector.

Figure 1: HPGe gammagraphy made by Cs-137

The information coming from the radiography or gammagraphy need to be blended with the data furnished by the device producer. The schemes furnished by producers are often poor in some structural detail. Typically, it is noticeable the absence of small internal clamps which, instead, are visible through a radiography or gammagraphy (see figure 1).

The second step deals with the proper simulation of the materials involved. In a Marinelli’s beaker must be defined, as accurately as possible, all the walls (typically made up of polypropylene or polystyrene) due to their thickness from 2 to 3 mm and, of course, the inner matrix. The matrix containing the radionuclides is well known both in density and in chemical composition in case of a calibration source being used. Almost always, in a typical situation, the matrix composition and density are more or less roughly known. As obvious, to overcome such a problems, it is necessary to setup several physical and chemical analysis to better understand the sample.

Materials existing only in trace, or constituting thin layers, can be neglected without altering, appreciably, the results. This is a great simplification in defining the matrix composition and it allows, also, to neglect thin layers inside HPGe endcup (i.e.: lithium or boron contact layers…). This simulation foresees only aluminum, germanium and vacuum to describe the HPGe detector.

The third step is to properly consider the dead layer (DL) of the HPGe crystal. It is fundamental to start the simulations from the nominal values declared by the producer. The dead layer definition is very important, in particular, in the medium-low energy range. About below 200 keV, small variations in DL give big variations in total efficiency simulations. However, the DL thicknesses give a further degree of freedom, in fact, enhancing the DL thickness it is possible to simulate an ageing of the crystal (i.e.: a general reduction in total efficiency). Moreover, adjusting DL offers the opportunity to overcome errors in determining the real thicknesses of the structures inside the endcup as they were estimated from the radiography or gammagraphy. The DL should be considered correct, in practice, if the simulated total efficiency lays within a proper range, such as 10% or less.

### The “source.f” file

Here it follows the source.f file for the specific case:

*$ CREATE SOURCE.FOR

*COPY SOURCE

*

*Elenco whats(x)

*Energia(MeV) - Zup - Zdown - Zrif - Rin - Rout // sdum: MAR

*Energia(MeV) - Zup - Zdown - R                 // sdum: CIL

*

*

*=== source Marinelli e Cilindrica ====================================*

*

SUBROUTINE SOURCE (NOMORE)

INCLUDE '(DBLPRC)'

INCLUDE '(DIMPAR)'

INCLUDE '(IOUNIT)'

*

*

INCLUDE '(BEAMCM)'

INCLUDE '(FHEAVY)'

INCLUDE '(FLKSTK)'

INCLUDE '(IOIOCM)'

INCLUDE '(LTCLCM)'

INCLUDE '(PAPROP)'

INCLUDE '(SOURCM)'

INCLUDE '(SUMCOU)'

*

LOGICAL LFIRST

DOUBLE PRECISION  Vup

DOUBLE PRECISION  Vdown

DOUBLE PRECISION  MYTETA

DOUBLE PRECISION  MYRHO

*

SAVE LFIRST

DATA LFIRST / .TRUE. /

*======================================================================*

*                 VERSION  1.0                                         *

*======================================================================*

NOMORE = 0

*  +-------------------------------------------------------------------*

*  |  First call initializations:

IF ( LFIRST ) THEN

*  |  *** The following 3 cards are mandatory ***

TKESUM = ZERZER

LFIRST = .FALSE.

LUSSRC = .TRUE.

*  |  *** User initialization ***

END IF

*  |

*  +-------------------------------------------------------------------*

*  Push one source particle to the stack. Note that you could as well

*  push many but this way we reserve a maximum amount of space in the

*  stack for the secondaries to be generated

* NPFLKA is the stack counter: of course any time source is called it

* must be =0

NPFLKA = NPFLKA + 1

* WTFLK is the weight of the particle

WTFLK     (NPFLKA) = ONEONE

WEIPRI = WEIPRI + WTFLK (NPFLKA)

* Particle type (1=proton 7=photon). Is the type set by the BEAM

* card

ILOFLK    (NPFLKA) = 7

* From this point .....

* Particle generation (1 for primaries)

LOFLK     (NPFLKA) = 1

* User dependent flag:

LOUSE  (NPFLKA) = 0

* User dependent spare variables:

DO 100 ISPR = 1, MKBMX1

SPAREK (ISPR,NPFLKA) = ZERZER

100  CONTINUE

* User dependent spare flags:

DO 200 ISPR = 1, MKBMX2

ISPARK (ISPR,NPFLKA) = 0

200  CONTINUE

* Save the track number of the stack particle:

ISPARK (MKBMX2,NPFLKA) = NPFLKA

NPARMA = NPARMA + 1

NUMPAR (NPFLKA) = NPARMA

NEVENT (NPFLKA) = 0

DFNEAR (NPFLKA) = +ZERZER

* ... to this point: don't change anything

* Particle age (s)

AGESTK (NPFLKA) = +ZERZER

AKNSHR (NPFLKA) = -TWOTWO

* Group number for "low" energy neutrons, set to 0 anyway

IGROUP (NPFLKA) = 0

* Kinetic energy of the particle (GeV)

TKEFLK  (NPFLKA) = WHASOU(1)*1.D-3

* Particle momentum

PMOFLK (NPFLKA) = SQRT ( TKEFLK (NPFLKA)

&       * ( TKEFLK (NPFLKA) + TWOTWO

&       * (AM (ILOFLK (NPFLKA) ) )) )

* Cosines (TXFLK,TYFLK,TZFLK)

CALL RACO(TXX,TYY,TZZ)

TXFLK (NPFLKA) = TXX

TYFLK (NPFLKA) = TYY

TZFLK (NPFLKA) = TZZ

* Polarization cosines:

TXPOL (NPFLKA) = -TWOTWO

TYPOL (NPFLKA) = +ZERZER

TZPOL (NPFLKA) = +ZERZER

* Particle coordinates

*======================================================================*

* Caso Marinelli  SDUSOU = MAR

IF (SDUSOU .EQ. 'MAR') THEN

Vup=(WHASOU(2)-WHASOU(4))*TWOPIP*(WHASOU(6)**2)

Vdown=(WHASOU(4)-WHASOU(3))*TWOPIP*(WHASOU(6)**2-WHASOU(5)**2)

*

IF (FLRNDM(V) .GE. (Vup/(Vup+Vdown))) THEN

* estrazione dalla ciambella

ZFLK (NPFLKA) = WHASOU(3)+(WHASOU(4)-WHASOU(3))

&        *FLRNDM(ZFLK)

MYTETA = TWOPIP*FLRNDM(TETA)

MYRHO = SQRT(WHASOU(5)**2+(WHASOU(6)**2-WHASOU(5)**2)

&        *FLRNDM(RHO))

ELSE

* estrazione dal cilindro

ZFLK (NPFLKA) = WHASOU(4)+(WHASOU(2)-WHASOU(4))

&        *FLRNDM(ZFLK)

MYTETA = TWOPIP*FLRNDM(TETA)

MYRHO = WHASOU(6)*SQRT(FLRNDM(RHO))

ENDIF

XFLK (NPFLKA) = MYRHO*COS(MYTETA)

YFLK (NPFLKA) = MYRHO*SIN(MYTETA)

*======================================================================*

ELSE IF (SDUSOU .EQ. 'CIL') THEN

ZFLK (NPFLKA) = WHASOU(3)+(WHASOU(2)-WHASOU(3))

&        *FLRNDM(ZFLK)

MYRHO = WHASOU(4)*SQRT(FLRNDM(RHO))

MYTETA = TWOPIP*FLRNDM(TETA)

XFLK (NPFLKA) = MYRHO*COS(MYTETA)

YFLK (NPFLKA) = MYRHO*SIN(MYTETA)

ELSE

ENDIF

*======================================================================*

*  Calculate the total kinetic energy of the primaries: don't change

IF ( ILOFLK(NPFLKA) .NE. 0 ) THEN

TKESUM = TKESUM + ( TKEFLK (NPFLKA) + AMDISC (ILOFLK(NPFLKA)) )

&          * WTFLK (NPFLKA)

ELSE

TKESUM = TKESUM + TKEFLK (NPFLKA) * WTFLK (NPFLKA)

END IF

*  Here we ask for the region number of the hitting point.

*     NRGFLK (NPFLKA) = ...

*  The following line makes the starting region search much more

*  robust if particles are starting very close to a boundary:

CALL GEOCRS ( TXFLK (NPFLKA), TYFLK (NPFLKA), TZFLK (NPFLKA) )

CALL GEOREG ( XFLK (NPFLKA), YFLK (NPFLKA), ZFLK (NPFLKA),

&              NRGFLK (NPFLKA), IDISC )

*  Do not change these cards:

CALL GEOHSM ( NHSPNT (NPFLKA), 1, -11, MLATTC )

NLATTC(NPFLKA) = MLATTC

CALL SOEVSV

RETURN

*=== End of subroutine Source =========================================*

END

The source is twofold, it can manage a Marinelli’s beaker and a cylinder source specifying “MAR” or “CIL” in the main “.inp” file. The event generator for the cylinder source is self explaining, whereas, in the case of Marinelli’s beaker it should be noticed the procedure that choose whether extracting from the upper portion of the beaker or in the lower one. The choice is made according to the proportion between the upper and lower decaying masses to obtain a proper probability distribution.

The source accept only one photon energy at each run.

The figure 2 displays a test to verify the proper events distribution inside a Marinelli’s beaker.

Figure 2: 3000 events distributed inside a Marinelli's beaker, as derived from the "source.f"

### The “.inp” file

Here it follows the “.inp” file in its most general format and for an Am-241 photon emission:

TITLE

Analisi in efficienza rivelatore HPGe - lab. radioprotezione

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

* ATTIVAZIONE RUN TOTALMENTE ANALOGICO

*GLOBAL        1000.0      -1.0       0.0       0.0       0.0       0.0

DEFAULTS                                                              PRECISIO

*

*

*inizio stesura:  30/05/2005

*ultima modifica: 30/06/2005

*

*NOTE: analisi gamma geometria MARINELLI

*NOTE: DeadLayer superiore di 1.5mm

*NOTE: DeadLayer esterno di 1.2mm circa

*NOTE: DeadLayer interno di 0.05mm circa

*============  GEOMETRIA INIZIO ============================================

GEOBEGIN                                                              COMBINAT

0    0           Detectors in HPGe

*

*asse Z in su - asse Y a destra - asse X entrante

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

* piani // XY HPGe

XYP    1       0.0

XYP    2       0.3

XYP    3       0.9

XYP    4       1.2

XYP    5       2.0

XYP    6       2.7

XYP    7       3.4

XYP    8       6.9

XYP    9       7.5

XYP   10       8.1

XYP   11       8.6

XYP   12      8.75

XYP   13      9.15

XYP   14       9.3

* cilindri HPGe

ZCC   15       0.0       0.0      0.15

ZCC   16       0.0       0.0     0.805

ZCC   17       0.0       0.0       1.0

ZCC   18       0.0       0.0      2.85

ZCC   19       0.0       0.0      2.95

ZCC   20       0.0       0.0      3.05

ZCC   21       0.0       0.0      3.45

ZCC   22       0.0       0.0       3.6

* sfera HPGe

SPH   23       0.0       0.0       6.9     0.805

* tronco di cono HPGe

TRC   24       0.0       0.0       8.1       0.0       0.0       0.5

2.85       2.5

* piani // XY Marinelli

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

XYP   25       1.8

XYP   26      2.05

XYP   27       9.3

XYP   28      9.55

XYP   29     16.60

XYP   30     16.65

XYP   31      16.9

* cilindri Marinelli

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

ZCC   32       0.0       0.0       3.6

ZCC   33       0.0       0.0      3.85

ZCC   34       0.0       0.0      5.70

ZCC   35       0.0       0.0      5.95

* cilindro e piani del limite geometria

RPP   36  -999999.   999999.  -999999.   999999.  -999999.   999999.

* cilindro deadlayer esterno 1mm per lato

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

ZCC   37       0.0       0.0      2.73

TRC   38       0.0       0.0       8.1       0.0       0.0      0.35

2.70      2.35

* deadlayer interno

SPH   39       0.0       0.0       6.9       0.8

ZCC   40       0.0       0.0       0.8

*

END

*

*============  DEFINIZIONE DELLE REGIONI  ==================================

*  SINTASSI 2X A3 I5 9(A2 I5)

*.000.....--.....--.....--.....--.....--.....--.....--.....--.....--.....

* buco nero esterno

001     OR +36    +31     -1    -35

002     OR +36    -31

003     OR +36     +1

* aria tra Marinelli e HPGe

004     OR +14     -1    -22    +32

* aria sotto Marinelli

005     OR +25     -1    -32    +35

* tappo Marinelli (POLISTIRENE)

006     OR +34    +31    -30

* fondo basso marinelli (PS)

007     OR +26    -25    +34    -33

* pareti esterne Marinelli (PS)

008     OR +31    -25    +35    -34

* pareti interne Marinelli (PS)

009     OR +28    -25    +33    -32

* fondo alto Marinelli (PS)

010     OR +28    -27    +32

*.000.....--.....--.....--.....--.....--.....--.....--.....--.....--.....

*============== interno del Marinelli =====================

* zona superiore vuota

011     OR +30    -29    +34

* >>>> volume di 1L del Marinelli <<<<

012     OR +29    -28    +34 OR +28    -26    +34    -33

*==========================================================

*==================== HPGe ================================

* HPGe

013     OR +38 OR +8     -5     +37    -16 OR +10    -8     +37    -23

*==========================================================

*.000.....--.....--.....--.....--.....--.....--.....--.....--.....--.....

* alluminio della capsula esterna protettiva HPGe (laterale)

014     OR +14     -1    +22    -21

* alluminio della capsula esterna protettiva HPGe (superiore)

015     OR +14    -13    +21

* alluminio della capsula interna protettiva HPGe (laterale)

016     OR +12     -1    +19    -18

* alluminio della capsula interna protettiva HPGe (superiore)

017     OR +12    -11    +18

* fascia superiore HPGe (alluminio)

018     OR  +9     -8    +20    -19

* fascia inferiore HPGe (alluminio)

019     OR  +7     -6    +20    -19

* stelo conduttore interno HPGe (alluminio)

020     OR  +8     -4    +15

* base piccola sotto lo stelo (alluminio)

021     OR  +4     -3    +40

* base media   sotto lo stelo (alluminio)

022     OR  +3     -2    +17

* base grande  sotto lo stelo (alluminio)

023     OR  +2     -1    +18

*.000.....--.....--.....--.....--.....--.....--.....--.....--.....--.....

* vuoto tra i rivestmenti parte alta

024     OR +13    -12    +19

* vuoto parte laterale sopra fascia superiore

025     OR +13     -9    +21    -19

* vuoto parte laterale accanto fascia superiore

026     OR  +9     -8    +21    -20

* vuoto parte laterale tra le due fasce

027     OR  +8     -7    +21    -19

* vuoto parte laterale accanto fascia inferiore

028     OR  +7     -6    +21    -20

* vuoto parte laterale sotto fascia inferiore

029     OR  +6     -1    +21    -19

* vuoto dentro sfera

030     OR +39     -8

* vuoto attorno stelo

031     OR  +8     -4    +40    -15

* vuoto attorno stelo sopra base media

032     OR  +5     -3    +18    -40

* vuoto attorno base meia sopra base grande

033     OR  +3     -2    +18    -17

* vuoto attrono tronco cono HPGe

034     OR +11    -10    +18    -24

*.000.....--.....--.....--.....--.....--.....--.....--.....--.....--.....

* dead layer esterno

035     OR -38    +24

036     OR +18    -37    +10     -5

* dead layer interno

037     OR +23    -39     -8

038     OR  +8     -5    +16    -40

*

END

GEOEND

*

*============  DEBUG DELLA GEOMETRIA  ================================

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

*GEOEND          30.0      30.0      30.0      -5.5      -5.5      -5.5DEBUG

*GEOEND         300.0     300.0     300.0                               &

*STOP

*

*GEOEND          10.0      10.0      10.0       0.0       0.0       0.0DEBUG

*GEOEND         100.0     100.0     100.0                               &

*STOP

*

*============  MATERIALI  ============================================

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

*MATERIAL         0.0       0.0       0.0       2.0                    VACUUM

*MATERIAL         1.0    1.0079  0.000089       3.0                    HYDROGEN

*MATERIAL         6.0    12.011       2.0       6.0                    CARBON

*MATERIAL         7.0   14.0067 0.0012506       7.0                    NITROGEN

*MATERIAL         8.0   15.9994  0.001429       8.0                    OXYGEN

*MATERIAL        12.0    24.305     1.738       9.0                    MAGNESIU

*MATERIAL        13.0    26.982    2.6989      10.0                    ALUMINUM

*MATERIAL        26.0    55.847     7.874      11.0                    IRON

*

MATERIAL        18.0    39.948  0.001784      26.0       0.0       0.0ARGON

MATERIAL        32.0     72.16      5.33      27.0       0.0       0.0GERMANIU

MATERIAL         0.0       0.0  0.001225      28.0       0.0       0.0AIR

MATERIAL         0.0       0.0      1.06      29.0       0.0       0.0POLIST

MATERIAL         0.0       0.0     1.394      30.0       0.0       0.0MYLAR

MATERIAL         0.0       0.0       1.0      31.0       0.0       0.0H2O

MATERIAL         0.0       0.0       0.9      32.0       0.0       0.0POLIPR

MATERIAL         0.0       0.0     1.032      33.0       0.0       0.0EPOXY

*

* definizione composti

COMPOUND       -0.78      -7.0     -0.21      -8.0     -0.01     -26.0AIR

COMPOUND         8.0       3.0       8.0       6.0                    POLIST

COMPOUND         8.0       3.0       3.0       6.0                    POLIPR

COMPOUND        19.0       3.0      18.0       6.0       3.0       8.0EPOXY

COMPOUND         8.0       3.0      10.0       6.0       4.0       8.0MYLAR

COMPOUND         2.0       3.0       1.0       8.0                    H2O

*

*============  ASSEGNAZIONE DEI MATERIALI ALLE REGIONI  ===============

*            materiale -> regione

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

ASSIGNMA          1.        1.        3.

ASSIGNMA         28.        4.        5.

ASSIGNMA         29.        6.       10.

*============ materiale Nel Marinelli ==============

* aria sopra pelo libero

ASSIGNMA         28.       11.

* materiale di riempimento

ASSIGNMA         33.       12.

*===================================================

*==================  HPGe  =========================

ASSIGNMA         27.       13.

*===================================================

ASSIGNMA         10.       14.       23.

ASSIGNMA          2.       24.       34.

*=========== HPGe deadlayer esterno ================

ASSIGNMA         27.       35.       36.

*=========== HPGe deadlayer interno ================

ASSIGNMA         27.       37.       38.

*

*============  DEFINIZIONE DELLA SORGENTE GAMMA  ======================

* elenco whats(x):

* Energia(MeV) - Zup - Zdown - Zrif - Rin - Rout // sdum: MAR

* Energia(MeV) - Zup - Zdown - R                 // sdum: CIL

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

* americio riga a 59.54 keV

SOURCE       0.05954     14.55      2.05      9.55      3.85       5.7MAR

*============  EMF E SOGLIE  ==========================================

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

EMF

EMFCUT       512.E-6     1.E-6       0.0       2.0      33.0       1.0PROD-CUT

*

EVENTYPE                                                              EVAP

PHYSICS          2.0                                                  EVAPORAT

*

*============  REGISTRAZIONE DEGLI EVENTI  ==============================

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

* rilevamento da 0 1024 keV

DETECT           0.0       0.0  1.024E-3                 0.0      13.0HPGE

* rilevamento da 1000 keV a 2024 keV

*DETECT           0.0     1.E-3  2.024E-3                 0.0      13.0HPGE

*====================================================================

*...+....1....+....2....+....3....+....4....+....5....+....6....+....7....+...

RANDOMIZE         1.

*

*start

*

START          2.E+7 99999999.                               99999000.

STOP

A typical result, “plotting a DETECT card”, can be seen in figure 3 for a Cs-137 gamma:

Figure 3: Cs-137 spectrum as simulated by FLUKA

The spectra displayed in figure 3 derives from simulating a Marinelli’s source with 661.6 keV photons from Cs-137. The matrix was made by epoxy resin with density of 1.0 g/cm^3.

The Dls were set to 1.5mm for the front-end of the device and to 1.2mm for the external surface.

The results of the performed simulations are displayed in figure 4:

Figure 4: experimental and simulated total efficiencies

The captions “SILENA” and “1,5+1,2+0,05mm” in figure 4 represents the experimental values and the simulated ones. The differences between experimental and simulated values lies in the range 10% as shown in table 1.

Table 1: comparison between experimental and simulated efficiencies

| E [keV] | Nuclide | Experimental values |  | DLs: 1.5 + 1.2 + 0.05 mm – FLUKA simulations |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | efficiency | uncertainty | efficiency | uncertainty | Difference % |
| 59,54 | Am-241 | 0,0071 | 0,0003 | 0,0072 | 0,0009 | 1,18% |
| 88,03 | Cd-109 | 0,0259 | 0,0009 | 0,0245 | 0,001 | -5,23% |
| 122,06 | Co-57 | 0,0363 | 0,0012 | 0,0351 | 0,0006 | -3,40% |
| 159,00 | Te-123m | 0,0377 | 0,0012 | 0,0371 | 0,0007 | -1,64% |
| 320,08 | Cr-51 | 0,0244 | 0,0007 | 0,0250 | 0,0007 | 2,62% |
| 391,69 | Sn-113 | 0,0227 | 0,0007 | 0,0213 | 0,0007 | -6,13% |
| 514,01 | Sr-85 | 0,0180 | 0,0005 | 0,0171 | 0,0007 | -5,14% |
| 661,66 | Cs-137 | 0,0155 | 0,0005 | 0,0141 | 0,0011 | -9,15% |
| 898,04 | Y-88 | 0,0113 | 0,0003 | 0,0113 | 0,0015 | 0,15% |
| 1173,24 | Co-60 | 0,0093 | 0,0003 | 0,0093 | 0,0007 | 0,02% |
| 1332,50 | Co-60 | 0,0083 | 0,0003 | 0,0084 | 0,0037 | 1,00% |
| 1836,06 | Y-88 | 0,0065 | 0,0002 | 0,0065 | 0,0063 | 0,15% |

Another check was performed with a thin disc source of Eu-152. The results shown in table 2, again, confirm the goodness of the simulation within 10%:

Table 2: Eu-152 disc source

|  emission | Absolute efficiency evaluation for an Eu-152 disc source reference |  |  |
| --- | --- | --- | --- |
| E [keV] | Experimental | DLs at 1.5+1.2+0.05 mm | Difference % |
| 121.78 | 0.14190.0004 | 0.15110.0001 | +6.5% |
| 344.28 | 0.08540.0003 | 0.08490.0003 | -0.6% |

### Conclusions

Tables 1 and 2 confirm the possibility to simulate an HPGe detector with FLUKA code. As obvious, the accurate knowledge of geometries and materials is crucial to obtain results with high accuracy. The principal difficulty is to determine the composition of the sample matrix. Anyway, accuracy within 10-20% can be obtained without excessive troubles.

### Further examples

ACTIVE CARBONS

The cylinder source (SDUM= CIL) undergoes further tests. The first test dealt with “active carbons” in the form of tiny powder. A polystyrene 100cc cylinder was filled with 42.262 g of active carbon and was measured for 72 h. After background subtraction, there were determined the CPS, as fitted by the acquisition software, of  Ra-226 and Th-232. More precisely, there were analyzed 3 peaks for each nuclide as follows:

| Chain | Energy | Nuclide |
| --- | --- | --- |
|  | keV |  |
| Ra-226 | 352 | Pb-214 |
| Th-232 | 583 | Tl-208 |
| Ra-226 | 609 | Bi-214 |
| Th-232 | 911 | Ac-228 |
| Ra-226 | 1120 | Bi-214 |
| Th-232 | 2614 | Tl-208 |

The cylinder source was filled in FLUKA with pure carbon at a density of  0.448 g/cm^3 and were performed 5 runs with 1*10^7 – 1*10^8 photons/run for each energy to collect the photopeak efficiency. Computing properly the branches, it was possible to collect the following values:

| Ra-226 | 68  10 Bq/kg |
| --- | --- |
| Th-232 | 33  7 Bq/kg |

The same sample was analyzed by another independent laboratory with different techniques and it was found that, again, the values were:

| Ra-226 | 68  6 Bq/kg |
| --- | --- |
| Th-232 | 35  5 Bq/kg |

The FLUKA simulation both of the HPGe and, above all, of the sample matrix were quite good.

### SimpleGeo (beta version 1.0) layout of detector interiors

| Gouraud Shading view | ZX-plane clip view |
| --- | --- |
| Wireframe view | Clip view |

Alessandro.porta@polimi.it