## Slide 1: Voxels and Medical Applications
FLUKA Advanced course

## Slide 2: The FLUKA voxel geometry
It is possible to describe a geometry in terms of “voxels”, i.e., tiny parallelepipeds (all of equal size) forming a 3-dimensional grid

## Slide 3: Voxel geometries: examples
The Gran Sasso in FLUKA

## Slide 4: Voxel geometries: examples
The anthropomorphic GOLEM phantom
Implementation in FLUKA
(radioprotection applications)
Petoussi-Henss et al, 2002
Notes: 4

## Slide 5: Voxel geometries in medical applications
Voxel geometries are especially useful to import CT scan of a human body, e.g., for dosimetric calculations of the planned treatment in radiotherapy
mGy
mGy
Parodi et al., 2007
FLUKA
Commercial TPS
Notes: 5

## Slide 6: Voxel geometry with PET-CT

## Slide 7: The FLUKA voxel geometry
The CT scan contains integer values “Hounsfield Unit” reflecting the X-ray attenuation coefficient mx
HUx = 1000 (mx-mH20) / mH20 , typically -1000  HU  3500
We will use loosely the word “organ” to indicate a group of voxels (or even more than one group) made of the same “tissue” material (same HU value or in a given HU interval)
The code handles each organ as a CG region, possibly in addition to other conventional “non-voxel” regions defined by the user
The voxel structure can be complemented by parts written in the standard Combinatorial geometry
The code assumes that the voxel structure is contained in a parallelepiped. This RPP is automatically generated from the voxel information.
In the past conversion programs customized by the user were needed, recently for medical applications FLAIR takes care of the conversion.
Notes: 7

## Slide 8: The FLUKA voxel geometry
To describe a voxel geometry, the user must convert his CT scan or equivalent data to a format understood by FLUKA. Starting from DICOM images, this is performed directly by FLAIR.
This stage should :
  Assign an organ index to each voxel. In many practical cases, the user will have a continuum of HU (CT values), and may have to group these values in intervals.
  Each organ is identified by a unique integer ≤32767. The organ numbering does not need to be contiguous i.e. “holes” in the numbering sequence are allowed.
  One of the organs must have number 0 and plays the role of the medium surrounding the voxels (usually vacuum or air).
  The user assigns to each NONZERO organ a voxel-region number. The voxel-region numbering has to be contiguous and starts from 1.

## Slide 9: The FLUKA voxel geometry
The information is input to FLUKA through a special file *.vxl containing:
  The number of voxels in each coordinate axis
  The voxel dimension in each coordinate axis
  The number of voxel-regions, and the maximum organ number
  A list of the organ corresponding to each voxel
  A list of the voxel-region number corresponding to each organ

## Slide 10: Input file
Prepare the usual FLUKA input file.
The geometry is written like a normal Combinatorial Geometry input, but in addition a VOXELS card must be inserted right after the GEOBEGIN card and before the Geometry title card
  WHAT(1), WHAT(2), WHAT(3) = x, y, z coordinates chosen as the origin of the “voxel volume”, (i.e. of a region made of a single RPP body extending from WHAT(1) to WHAT(1) + NX*DX, …) which contains all the voxels
  WHAT(4) ROT-DEFI transformation applied to whole voxel RPP
  WHAT(5), WHAT(6): not used
  SDUM = name of the voxel file
  extension will be assumed to be .vxl)

## Slide 11: Processing the DICOM files with FLAIR I
DICOM = Digital Imaging and Communications in Medicine is a medical standard for distributing any kind of medical image.
FLAIR has a capability to process the DICOM files using the pydicom module and convert them to FLUKA VOXELS or USRBIN compatible files.
First select the “Directory” where the DICOM data sets are located (if you have doubts press F1 and the flair manual will help you).
Select one “Data sets” and inspect the images.

## Slide 12: Processing the DICOM files with FLAIR II
You can crop the voxel by graphically selecting the region of interest

## Slide 13: Processing the DICOM files with FLAIR III

## Slide 14: Processing the DICOM files with FLAIR IV

## Slide 15: Processing the DICOM files with FLAIR V
The “Voxel” tab is used in order to convert the dataset to VOXELS or USRBIN format. For the VOXEL geometry two additional files are needed (example: material.inp and head.mat in the flair/dicom dir).

## Slide 16: Processing the DICOM files with FLAIR VI
≤ Unit: specify the upper limit of the range. Every entry will correspond to a range from the previous upper limit+1 until the current upper limit.
Material: select any of the predefined FLUKA materials defined previously.
Optionally you can specify correction factors for the density and dE/dx
Crho_min/Chro_max: density correction factors to be applied on the lower/upper limit of the unit range (see next slides).
CdEdx_rel_min/CdEdx_rel_max: relative correction factors on dE/dx for minimum/maximum unit in the range (see next slides).

## Slide 17: Voxel Body
The usual list of NB bodies, not including the RPP corresponding to the “voxel volume” (see VOXELS card above). This RPP will be generated and added automatically by the code as the (NB+1) th body, with one corner in the point indicated in the VOXELS card, and dimensions NX*DX, NY*DY and NZ*DZ as read from the voxel file.
The usual region list of NR regions, with the space occupied by body named VOXEL or numbered NB+1 (the “voxel volume”) subtracted. In other words, the NR regions listed must cover the whole available space, excepted the space corresponding to the “voxel volume”. This is easily obtained by subtracting body VOXEL or NB+1 in the relevant region definitions, even though this body is not explicitly input at the end of the body list.

## Slide 18: Voxel Regions
The code will automatically generate and add several regions:
NO additional regions, where NO = number of non-zero organs:
Name | Number | Description
VOXEL | NR+1 | sort of a “cage” for all voxels. Nothing should ever be deposited in it. The user shall assign vacuum to it.
VOXEL001 | NR+2 | containing all voxels belonging to organ number 0. There must be at least 2 of such voxels, but in general they should be many more. Typical material assignment to this region is air
VOXEL002 | NR+3 | corresponding to organ 1
VOXEL003 | NR+4 | corresponding to organ 2
VOXEL### | NR+2+NO | corresponding to organ NO

## Slide 19: Few remarks
The assignment of materials is made directly by FLAIR. The user has to assign the materials to the regions defined by combinatorial geometry.
The “head.mat” and “material.inp” files are examples, the user should update these files taking into account his calibration curves.

## Slide 20: Practical issues for Medical Applications
General problems for MC calculations on CT scans
How to assign realistic human tissue parameters (= materials) for MC Calculation?
How to find a good compromise between the number of different HU values (~ 3000-5000) and the materials to be considered in the MC?
(issues on memory and computation speed when attempting to treat each HU number as a different material !!!)
How to preserve continuous, HU-dependent information when segmenting the HU numbers into intervals sharing the same “tissue” material ?
(critical for ion range calculation in hadron therapy !!!)
Notes: 20

## Slide 21: CT stoichiometric calibration (I)
CT segmentation into 27 materials of defined elemental composition (from analysis of 71 human CT scans)
Soft tissue
Air, Lung, Adipose tissue
Skeletal tissue
Schneider et al PMB 45, 2000
Notes: 21

## Slide 22: CT stoichiometric calibration (II)
Assign to each material a “nominal mean density”, e.g. using the density at the center of each HU interval (Jiang et al, MP 2004)
Schneider et al PMB 45, 2000
But “real density” (and related physical quantities) varies continuously with HU value !!!

## Slide 23: The region-dependent CORRFACT card
“CORRFACT” card allows to alter material density for dE/dx and nuclear processes
First two inputs specify a density scaling factor (restricted to the interval [2/3,3/2]) for charged particle ionization processes (WHAT(1)) and for all other processes (WHAT(2)) to the region(s) specified by the inputs WHAT(4-6) [cf. manual]
This is especially important in ion beam therapy to force the MC to follow the same semi-empirical HU-range calibration curve as the Treatment Planning System (TPS) for dosimetric comparisons
FLAIR automatically appends the CORRFACT cards calculated taking into account the calibration curves provided by the user at the end of the .vxl file.

## Slide 24: How to account for HU-dependent dEdx
In the INPUT
  Let several regions share the same material composition and mean density according to CT segmentation (reduced number of materials to save memory / initialization time)
  ASSIGNMA BONE VOXEL005 (region number 25)
  ASSIGNMA BONE VOXEL016 (region number 31)
  Use CORRFACT to impose the desired correction for stopping power ( ion range!) in the regions KREG corresponding to different organs IO (i.e., different HU values) sharing the same MATERIAL assignment
  CORRFACT 0.85 0.0 0.0 VOXEL005
  CORRFACT 1.3 0.0 0.0 VOXEL016
Region #25 corresponds
to “softer” bone than #31

## Slide 25: Forcing FLUKA to follow the same range calibration curve as TPS for p @ MGH Boston
Parodi et al MP 34, 2007, Parodi et PMB 52, 2007
1000
The CORRFACT ionization scaling factors were obtained from the dEdx ratio between TPS and FLUKA (+ Schneider “mass density”) -> The user should update the “head.mat” file with his own calibration for CdEdX_rel taking into account his TPS.
Schneider’s segmentation
24 Materials, 2600 regions (HU values)
Ti
Extrapolation to Ti
3 Materials, 1900 regions
(HU values)

## Slide 26: Applications of FLUKA to p therapy @ MGH
TPS
FLUKA
Prescribed dose: 1 GyE
MC : ~ 5.5 106 protons in 10 independent runs (11h each on Linux Cluster mostly using 2.2GHz Athlon processors)
Parodi et al, JPCS 74, 2007
mGy
mGy
Input phase-space provided by H. Paganetti, MGH Boston

## Slide 27: Applications of FLUKA to p therapy @ MGH
TPS
FLUKA
mGy
mGy
FLUKA
TPS
mGy
mGy
Parodi et PMB 52, 2007

## Slide 28
FLUKA
mGy
Prescribed dose: 2 GyE
MC : ~ 7.4 107p in 12 independent runs (~ 130h each on 2.2 GHz Linux cluster)
TPS
metallic
implants
K. Parodi et al, IJROBP 2007
mGy
Applications of FLUKA to p therapy @ MGH

## Slide 29: Biological predictions with mixed field
Estimation of the biological effect is important in many applications (hadron therapy, space radiation, etc…)
The linear-quadratic model for the biological response predictions is often used :
S = exp (-α * D – β * D *D)
where S is the (cell, tissue) survival and α and β are the linear and quadratic term for the dose response.
α and β values depends from many factors:
  Particle type (Z, A)
  Particle energy
  Cell / Tissue type
  Biological endpoint
  …
α and β values are passed to FLUKA with RAD-BIOL card

## Slide 30: RAD-BIOL
RAD-BIOL: reading biological parameter cards
  WHAT(1) = 0 always
  WHAT(2-6): not used
  SDUM = biological data file name (10 character long) the file must be named <sdum>.dat
Example of an input biological table for H, He and C ions expressed
as RBEα = αion / αphotons as function of MeV /u

## Slide 31: TPSSCORE
This card controls some scoring parameters relevant for TPS-like calculations.
  For SDUM = BIOTOBIN
  WHAT(1) = > 0, alpha-beta cellular line index to be associated to the USRBIN defined by WHAT(4)-WHAT(6) (see RAD-BIOL)
  WHAT(2), WHAT(3) = not used
  WHAT(4) = from binning WHAT(4)
  WHAT(5) = to binning WHAT(5)
  WHAT(6) = in step of WHAT(6)
RAD-BIOL TABLENAMAB USRBIN 11. ALPHA-D -42. 10. 0. 30.0ALPHADa USRBIN 0. 0. 0.0 100. 600.0 & USRBIN 11. SQBETA-D -42. 10. 0. 30.0BETADa USRBIN 0. 0. 0.0 100. 600.0 & TPSSCORE 1. ALPHADa BETADa BIOTOBIN
Notes: For SDUM=BIOTOBIN WHAT(1) = > 0 alpha-beta cellular line index to be associated to the usrbin defined by WHAT(4)-WHAT(6) (see RAD-BIOL) < 0 resets to default (=0) = 0 ignored WHAT(2) = not used WHAT(3) = not used WHAT(4) = from binning "WHAT(4)" .... (def. 1) WHAT(5) = .... to binning "WHAT(5)" .... (def. WHAT(4)) WHAT(6) = .... in step of "WHAT(6)" (def. 1)

## Slide 32: TPSSCORE
Use TPSSCORE to calculate the mixed radiation field liner and quadratic term (left panel)
Use USRBIN dose (D) and TPSSCORE to calculate biological dose (DRBE, right panel in red) in the post-processing
G. Battistoni, et al. Frontiers in Oncology 2016

## Slide 33: FLUKA biological application for 12C ion therapy
A. Mairani, et al. PMB 2010
in vitro data
predictions
G. Magro, et al. PMB 2017
FLUKA + LEM model
FLUKA + NIRS approach

## Slide 34: Applications of FLUKA to p therapy @ MGH
mGy
mGy
Clival Chordoma, 0.96 GyE / field
Planned dose

## Slide 35: Post-radiation PET/CT @ MGH
Clival Chordoma, 0.96 GyE / field, DT1 ~ 26 min, DT2 ~ 16 min
Bq / ml
Bq / ml
Average Activity
1 Field
2 Field
K. Parodi et al, IJROBP 2007
… and FLUKA-voxel functionalities being also used at HIT and CNAO …